# Instrumentation Calibration and Maintenance Standards

## Introduction

Regular calibration and maintenance underpin the accuracy and reliability of refinery instrumentation. This document establishes procedures and intervals for calibrating instruments, managing spares and documenting maintenance activities.

## Applicable Standards

The following standards define requirements for calibration laboratories, management systems and maintenance practices:

- API 551 – Instrument Calibration – guidance on calibration procedures for refinery instruments.

- ISO 10012 – Measurement Management Systems – requirements for measurement processes and metrological confirmation.

- ISO/IEC 17025 / MIL‑STD‑45662A – general requirements for the competence of calibration and testing laboratories and calibration system requirements【257831869405972†L528-L540】.

- ANSI/ISA RP 12.6.01 – Maintenance and Testing of Fire & Gas Detectors – recommended practice for testing and maintaining fire and gas detection systems.

## Calibration Equipment

- Pressure calibrators: deadweight testers, pressure pumps and electronic calibrators with NIST‑traceable accuracy.

- Temperature calibrators: dry‑block calibrators, fluid baths, reference RTDs and thermocouples.

- Flow calibrators: primary standards (provers), portable ultrasonic or Coriolis calibrators.

- Electrical loop calibrators: simulate and measure 4–20 mA signals, voltage and resistance.

- Communicators: HART/Fieldbus communicators for configuration and diagnostics.

## Calibration Procedures

- Schedule – develop a calibration schedule based on criticality, manufacturer recommendations and regulatory requirements. For safety‑critical instruments, calibrate at least annually.

- Pre‑calibration check – inspect instrument condition; verify correct installation and absence of damage or leaks.

- Calibration points – calibrate at multiple points across the operating range (e.g., 0, 25, 50, 75 and 100 %); record as‑found and as‑left data.

- Adjustment – adjust zero, span and linearity using manufacturer procedures; minimise hysteresis.

- Documentation – record calibration data, date, technician, calibration equipment used and certificates; file in maintenance management system.

- Out‑of‑tolerance instruments – remove from service; tag and repair or replace; evaluate process impact and implement corrective action.

## Maintenance Practices

- Preventive maintenance – perform routine inspections (visual check of cables, enclosures, indications); clean sensors; replace consumables (filters, membranes, diaphragms); lubricate moving parts.

- Predictive maintenance – Leverage smart diagnostics from transmitters, analyzers and control valves to monitor health (sensor drift, valve signature, plugged impulse lines) and apply data analytics or artificial intelligence to predict failures; schedule maintenance based on these condition indicators to avoid unplanned downtime.

- Corrective maintenance – repair or replace failed instruments promptly; maintain spare parts inventory.

- Spare management – maintain a list of critical spares for transmitters, sensors, electronics, regulators and fittings; store in controlled environment.

- Calibration tools maintenance – calibrate and certify calibration equipment regularly; store in protective cases; charge batteries and update firmware.

## Safety Considerations

- Follow lockout/tagout procedures before removing instruments.

- Vent and depressurise impulse lines before disconnecting.

- Use appropriate personal protective equipment (PPE) when working on process equipment.

- For hazardous areas, ensure work permits and gas testing are completed.

## Training and Competence

- Calibration and maintenance shall be performed by qualified technicians trained on specific instrument types.

- Provide ongoing training on new technologies, safety regulations and maintenance software.

## Record Keeping

- Maintain calibration certificates and maintenance records for audit and compliance.

- Use computerised maintenance management systems (CMMS) or instrument asset management systems (AMS) for planning and tracking.